## Dashboard
Variation Analysis
Benchmark Run Results
Test Run Results
Reagents used
Concordance Study
Conducted validation testing on new lot numbers for Binding Solution and Binding Beads to ensure compliance with established protocols and consistency with previous results, prior to integration into routine laboratory use.
Results Concordance:
94.23%
Study Outcome:
Fail
MS2 Concordance:
7.53%
Test MS2 Standard Deviation:
0.88
Concluding Remarks
There is significant variation between the benchmark and test runs, as evident from the shift in MS2 Ct values by a variance of 5 Cts. This indicates that the test binding solution and binding beads are less efficient in facilitating PCR amplification, as evidenced by the delayed Ct values. This shift is causing some of the controls to remain undetected due to late amplification. Additional analysis is necessary to pinpoint whether the binding solution or the binding beads are responsible for this shift.
Test Reagents:
Binding Solution
Binding Beads

## Variance
Variation Analysis
|  |  | Benchmark - Run 1 |  |  |  |  | Test - Run 2 |  |  |  |  | Variance (<2.5)
Sample Type | Concentration (copies/ml) | Sample Name | MS2 Ct | FluAB Ct | RSV Ct | C19 Ct | Detection | MS2 Ct | FluAB Ct | RSV Ct | C19 Ct | Detection | MS2 Ct | FluAB Ct | RSV Ct | C19 Ct | Detection
Positive Control | 1:50 Dilution | PC_1 | Undetermined | 23.04 | 24.71 | 27.69 | FluAB+,RSV+,C19+ | Undetermined | 22.14 | 24.56 | 27.52 | FluAB+,RSV+,C19+ | Not Applicable | 0.9 | 0.15 | 0.17 | FluAB+,RSV+,C19+
Positive Control | 1:2,800 Dilution | PC_2 | Undetermined | 28.55 | 30.52 | 32.93 | FluAB+,RSV+,C19+ | Undetermined | 28.53 | 30.37 | 33 | FluAB+,RSV+,C19+ | Not Applicable | 0.02 | 0.15 | 0.07 | FluAB+,RSV+,C19+
Negative Control | 0 | NC_1 | 25.77 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 27.22 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 1.45 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Environmental Water Control | 0 | 1320496 | 25.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.25 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.99 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
High IQC | 10000 | 1320810 | 24.64 | Undetermined | Undetermined | 28.55 | C19+ | 29.76 | Undetermined | Undetermined | 33.94 | C19+ | 5.12 | Not Applicable | Not Applicable | 5.39 | C19+
High IQC | 10000 | 1320811 | 25.33 | Undetermined | Undetermined | 29.29 | C19+ | 26.77 | Undetermined | Undetermined | 34.31 | C19+ | 1.44 | Not Applicable | Not Applicable | 5.02 | C19+
High IQC | 10000 | 1320812 | 25.52 | Undetermined | Undetermined | 29.57 | C19+ | 29.75 | Undetermined | Undetermined | 34.5 | C19+ | 4.23 | Not Applicable | Not Applicable | 4.93 | C19+
Low IQC | 1000 | 1336368 | 25.49 | Undetermined | Undetermined | 32.84 | C19+ | 29.41 | Undetermined | Undetermined | 40.84 | FluAB,RSV,C19 Not Detected | 3.92 | Not Applicable | Not Applicable | 8 | Discrepancy
Low IQC | 1000 | 1336369 | 25.39 | Undetermined | Undetermined | 33.21 | C19+ | 29.25 | Undetermined | Undetermined | 40.49 | FluAB,RSV,C19 Not Detected | 3.86 | Not Applicable | Not Applicable | 7.28 | Discrepancy
Low IQC | 1000 | 1338669 | 25.06 | Undetermined | Undetermined | 32.65 | C19+ | 29.4 | Undetermined | Undetermined | 36.72 | C19+ | 4.34 | Not Applicable | Not Applicable | 4.07 | C19+
Low Low IQC | 500 | 1320804 | 25.15 | Undetermined | Undetermined | 33.79 | C19+ | 29.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.11 | Not Applicable | Not Applicable | Not Applicable | Discrepancy
Low Low IQC | 500 | 1320805 | 25.26 | Undetermined | Undetermined | 33.7 | C19+ | 29.61 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.35 | Not Applicable | Not Applicable | Not Applicable | Discrepancy
Low Low IQC | 500 | 1320806 | 25.11 | Undetermined | Undetermined | 33.81 | C19+ | 26.24 | Undetermined | Undetermined | 38.96 | FluAB,RSV,C19 Not Detected | 1.13 | Not Applicable | Not Applicable | 5.15 | Discrepancy
Negative IQC | 0 | 1320801 | 25.32 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.48 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative IQC | 0 | 1320802 | 25.35 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.02 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.67 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative IQC | 0 | 1320803 | 25.11 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.98 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.87 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Known Positive Patient Sample |  | MWA37090217 | 25.78 | Undetermined | Undetermined | 19.58 | C19+ | 30.32 | Undetermined | Undetermined | 24.39 | C19+ | 4.54 | Not Applicable | Not Applicable | 4.81 | C19+
Known Positive Patient Sample |  | MWA46425086 | 25.52 | Undetermined | Undetermined | 21.17 | C19+ | 30.79 | Undetermined | Undetermined | 27.69 | C19+ | 5.27 | Not Applicable | Not Applicable | 6.52 | C19+
Known Positive Patient Sample |  | VIV10319351 | 26.21 | Undetermined | Undetermined | 22.88 | C19+ | 31.68 | Undetermined | Undetermined | 27.74 | C19+ | 5.47 | Not Applicable | Not Applicable | 4.86 | C19+
Known Positive Patient Sample |  | DHA22279997 | 25.53 | Undetermined | Undetermined | 26.27 | C19+ | 28.81 | Undetermined | Undetermined | 29.23 | C19+ | 3.28 | Not Applicable | Not Applicable | 2.96 | C19+
Known Positive Patient Sample |  | MWA62885277 | 26.08 | Undetermined | Undetermined | 26.49 | C19+ | 31.63 | Undetermined | Undetermined | 32.37 | C19+ | 5.55 | Not Applicable | Not Applicable | 5.88 | C19+
Known Positive Patient Sample |  | DHA16965561 | 25.83 | Undetermined | Undetermined | 31.7 | C19+ | 29.71 | Undetermined | Undetermined | 43.09 | FluAB,RSV,C19 Not Detected | 3.88 | Not Applicable | Not Applicable | 11.39 | Discrepancy
Known Positive Patient Sample |  | KNG62150659 | 25.79 | Undetermined | Undetermined | 27.31 | C19+ | 31.59 | Undetermined | Undetermined | 32.57 | C19+ | 5.8 | Not Applicable | Not Applicable | 5.26 | C19+
Known Positive Patient Sample |  | MWA36401816 | 25.31 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.4 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.09 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Known Positive Patient Sample |  | MWA48944684 | 26.28 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.08 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.8 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320742 | 25.14 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.74 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.6 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320743 | 25.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320746 | 25.38 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 31.2 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 5.82 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320749 | 25.52 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.56 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.04 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320750 | 25.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.56 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 5.11 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320751 | 25.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.96 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.7 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320752 | 25.34 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.95 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.61 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320753 | 25.38 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.01 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.63 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320756 | 25.42 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.8 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 5.38 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320757 | 25.42 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.39 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.97 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320758 | 25.31 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.17 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.86 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320759 | 25.2 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.06 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.86 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320761 | 25.49 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.13 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.64 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320762 | 25.29 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.27 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.98 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320763 | 25.0794 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.4166 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.3372 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320764 | 25.5678 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.4948 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.927 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320765 | 26.0169 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6216 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.6047 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320766 | 25.3251 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.5215 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1964 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320767 | 25.0672 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.2177 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 5.1505 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320768 | 25.5437 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8271 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.2834 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320769 | 25.5294 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6571 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1277 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320770 | 25.6847 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.9277 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.243 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320771 | 25.5445 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.5252 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.9807 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320772 | 25.3509 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.576 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.2251 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320773 | 25.8408 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.3265 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.4857 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320774 | 25.3504 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6212 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.2708 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320775 | 25.0887 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.3975 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.3088 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320776 | 25.8932 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.5152 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.622 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320777 | 25.4539 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6509 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1969 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320778 | 25.5037 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.7261 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.2224 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320779 | 25.2281 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8143 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5861 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320780 | 26.2709 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.7558 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.4848 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320781 | 26.0673 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.1505 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.0832 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320782 | 25.0471 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.6503 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.6031 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320783 | 25.0236 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.3692 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.3456 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320784 | 26.2096 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8712 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.6615 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320785 | 25.0439 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6075 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5636 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320786 | 25.7074 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.9842 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.2768 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320787 | 26.1002 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.5803 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.4801 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320788 | 25.0036 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.1725 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.1688 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320789 | 26.3676 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.9609 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5933 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320790 | 26.4342 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8836 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.4495 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320791 | 26.2078 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.8813 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.6735 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320792 | 25.0423 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.7521 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.7098 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320793 | 25.4566 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.5637 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1072 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320794 | 25.4679 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.7305 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.2626 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320795 | 26.3477 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.8797 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.532 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320796 | 26.3317 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.5278 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1961 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320797 | 25.533 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.8736 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.3406 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320798 | 26.483 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.9997 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5167 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320799 | 26.1488 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.004 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.8552 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320800 | 26.0116 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.9813 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.9696 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320807 | 26.4798 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.4802 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.0004 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320808 | 26.1106 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.419 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.3084 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320809 | 25.7302 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8749 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.1447 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320813 | 25.8403 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.2385 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.3982 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320814 | 25.997 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.5304 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5335 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320815 | 25.7527 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.6498 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.897 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320816 | 25.539 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.7016 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.1626 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320817 | 25.2278 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.0244 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.7967 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320818 | 25.3856 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.2254 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.8398 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320819 | 25.5541 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.25 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.6959 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320820 | 25.1264 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.114 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.9877 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320821 | 26.0384 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.6992 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.6608 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320822 | 26.4957 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.5247 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.029 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320823 | 25.4377 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 30.1316 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 4.694 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320824 | 25.0371 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.7374 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.7003 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320825 | 25.6991 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.5478 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.8487 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320826 | 26.1017 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 28.8929 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 2.7912 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
Negative Template Control | 0 | 1320827 | 25.5755 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 29.0977 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected | 3.5222 | Not Applicable | Not Applicable | Not Applicable | FluAB,RSV,C19 Not Detected
| Standard Deviation of MS2: |  | 0.4205 |  |  |  | Standard Deviation of MS2: | 0.8801 |  |  |  | MS2 Percentage Concordance: | 7.527% |  | Results Concordance: |  | 94.23%
[15 empty] | Study Outcome (>95%): |  | Fail

## Benchmark Run 1
Tag | Sample Name | MS2 Ct | FluAB Ct | RSV Ct | C19 Ct | Detection
Control | PC_1 | Undetermined | 23.04 | 24.71 | 27.69 | FluAB+,RSV+,C19+
Control | PC_2 | Undetermined | 28.55 | 30.52 | 32.93 | FluAB+,RSV+,C19+
Control | NC_1 | 25.77 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Env Water | 1320496 | 25.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
High | 1320810 | 24.64 | Undetermined | Undetermined | 28.55 | C19+
High | 1320811 | 25.33 | Undetermined | Undetermined | 29.29 | C19+
High | 1320812 | 25.52 | Undetermined | Undetermined | 29.57 | C19+
Low | 1336368 | 25.49 | Undetermined | Undetermined | 32.84 | C19+
Low | 1336369 | 25.39 | Undetermined | Undetermined | 33.21 | C19+
Low | 1338669 | 25.06 | Undetermined | Undetermined | 32.65 | C19+
Low Low | 1320804 | 25.15 | Undetermined | Undetermined | 33.79 | C19+
Low Low | 1320805 | 25.26 | Undetermined | Undetermined | 33.7 | C19+
Low Low | 1320806 | 25.11 | Undetermined | Undetermined | 33.81 | C19+
Neg | 1320801 | 25.32 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Neg | 1320802 | 25.35 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Neg | 1320803 | 25.11 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Strong Ct | MWA37090217 | 25.78 | Undetermined | Undetermined | 19.58 | C19+
Strong Ct | MWA46425086 | 25.52 | Undetermined | Undetermined | 21.17 | C19+
Strong Ct | VIV10319351 | 26.21 | Undetermined | Undetermined | 22.88 | C19+
Medium Ct | DHA22279997 | 25.53 | Undetermined | Undetermined | 26.27 | C19+
Medium Ct | MWA62885277 | 26.08 | Undetermined | Undetermined | 26.49 | C19+
Medium Ct | DHA16965561 | 25.83 | Undetermined | Undetermined | 31.7 | C19+
Weak Ct | KNG62150659 | 25.79 | Undetermined | Undetermined | 27.31 | C19+
Weak Ct | MWA36401816 | 25.31 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Weak Ct | MWA48944684 | 26.28 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320742 | 25.14 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320743 | 25.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320746 | 25.38 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320749 | 25.52 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320750 | 25.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320751 | 25.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320752 | 25.34 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320753 | 25.38 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320756 | 25.42 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320757 | 25.42 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320758 | 25.31 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320759 | 25.2 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320761 | 25.49 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320762 | 25.29 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320763 | 25.0794 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320764 | 25.5678 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320765 | 26.0169 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320766 | 25.3251 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320767 | 25.0672 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320768 | 25.5437 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320769 | 25.5294 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320770 | 25.6847 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320771 | 25.5445 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320772 | 25.3509 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320773 | 25.8408 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320774 | 25.3504 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320775 | 25.0887 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320776 | 25.8932 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320777 | 25.4539 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320778 | 25.5037 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320779 | 25.2281 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320780 | 26.2709 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320781 | 26.0673 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320782 | 25.0471 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320783 | 25.0236 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320784 | 26.2096 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320785 | 25.0439 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320786 | 25.7074 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320787 | 26.1002 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320788 | 25.0036 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320789 | 26.3676 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320790 | 26.4342 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320791 | 26.2078 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320792 | 25.0423 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320793 | 25.4566 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320794 | 25.4679 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320795 | 26.3477 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320796 | 26.3317 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320797 | 25.533 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320798 | 26.483 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320799 | 26.1488 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320800 | 26.0116 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320807 | 26.4798 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320808 | 26.1106 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320809 | 25.7302 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320813 | 25.8403 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320814 | 25.997 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320815 | 25.7527 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320816 | 25.539 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320817 | 25.2278 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320818 | 25.3856 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320819 | 25.5541 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320820 | 25.1264 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320821 | 26.0384 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320822 | 26.4957 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320823 | 25.4377 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320824 | 25.0371 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320825 | 25.6991 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320826 | 26.1017 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320827 | 25.5755 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected

## Test Run 2
Tag | Sample Name | MS2 Ct | FluAB Ct | RSV Ct | C19 Ct | Detection
Control | PC_1 | Undetermined | 22.14 | 24.56 | 27.52 | FluAB+,RSV+,C19+
Control | PC_2 | Undetermined | 28.53 | 30.37 | 33 | FluAB+,RSV+,C19+
Control | NC_1 | 27.22 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Env Water | 1320496 | 30.25 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
High | 1320810 | 29.76 | Undetermined | Undetermined | 33.94 | C19+
High | 1320811 | 26.77 | Undetermined | Undetermined | 34.31 | C19+
High | 1320812 | 29.75 | Undetermined | Undetermined | 34.5 | C19+
Low | 1336368 | 29.41 | Undetermined | Undetermined | 40.84 | FluAB,RSV,C19 Not Detected
Low | 1336369 | 29.25 | Undetermined | Undetermined | 40.49 | FluAB,RSV,C19 Not Detected
Low | 1338669 | 29.4 | Undetermined | Undetermined | 36.72 | C19+
Low Low | 1320804 | 29.26 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Low Low | 1320805 | 29.61 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Low Low | 1320806 | 26.24 | Undetermined | Undetermined | 38.96 | FluAB,RSV,C19 Not Detected
Neg | 1320801 | 28.8 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Neg | 1320802 | 29.02 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Neg | 1320803 | 28.98 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Strong Ct | MWA37090217 | 30.32 | Undetermined | Undetermined | 24.39 | C19+
Strong Ct | MWA46425086 | 30.79 | Undetermined | Undetermined | 27.69 | C19+
Strong Ct | VIV10319351 | 31.68 | Undetermined | Undetermined | 27.74 | C19+
Medium Ct | DHA22279997 | 28.81 | Undetermined | Undetermined | 29.23 | C19+
Medium Ct | MWA62885277 | 31.63 | Undetermined | Undetermined | 32.37 | C19+
Medium Ct | DHA16965561 | 29.71 | Undetermined | Undetermined | 43.09 | FluAB,RSV,C19 Not Detected
Weak Ct | KNG62150659 | 31.59 | Undetermined | Undetermined | 32.57 | C19+
Weak Ct | MWA36401816 | 28.4 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
Weak Ct | MWA48944684 | 30.08 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320742 | 28.74 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320743 | 28.45 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320746 | 31.2 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320749 | 28.56 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320750 | 30.56 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320751 | 28.96 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320752 | 28.95 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320753 | 30.01 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320756 | 30.8 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320757 | 30.39 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320758 | 30.17 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320759 | 30.06 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320761 | 30.13 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320762 | 29.27 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320763 | 28.4166 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320764 | 29.4948 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320765 | 28.6216 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320766 | 28.5215 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320767 | 30.2177 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320768 | 28.8271 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320769 | 28.6571 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320770 | 28.9277 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320771 | 29.5252 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320772 | 29.576 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320773 | 29.3265 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320774 | 28.6212 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320775 | 29.3975 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320776 | 29.5152 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320777 | 28.6509 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320778 | 29.7261 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320779 | 28.8143 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320780 | 29.7558 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320781 | 30.1505 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320782 | 29.6503 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320783 | 29.3692 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320784 | 28.8712 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320785 | 28.6075 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320786 | 28.9842 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320787 | 29.5803 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320788 | 29.1725 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320789 | 29.9609 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320790 | 28.8836 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320791 | 29.8813 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320792 | 29.7521 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320793 | 28.5637 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320794 | 28.7305 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320795 | 29.8797 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320796 | 29.5278 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320797 | 29.8736 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320798 | 29.9997 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320799 | 29.004 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320800 | 29.9813 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320807 | 28.4802 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320808 | 28.419 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320809 | 28.8749 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320813 | 29.2385 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320814 | 29.5304 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320815 | 29.6498 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320816 | 29.7016 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320817 | 30.0244 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320818 | 29.2254 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320819 | 29.25 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320820 | 29.114 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320821 | 28.6992 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320822 | 28.5247 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320823 | 30.1316 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320824 | 28.7374 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320825 | 28.5478 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320826 | 28.8929 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected
NTC | 1320827 | 29.0977 | Undetermined | Undetermined | Undetermined | FluAB,RSV,C19 Not Detected

## Reagents
Reagents used for Concordance Study
Reagents | Benchmark Run Lot numbers | Test Run Lot Numbers
COVID IQC | Old | Old
HBSS | Old | Old
Binding Solution | Old | New
Binding Beads | Old | New
MS2 | Old | Old
Proteinase K | Old | Old
HT water for NC | Old | Old
Wash 1 | Old | Old
Ethanol | Old | Old
Elution Solution | Old | Old
PC | Old | Old
Dilution Buffer | Old | Old
Mastermix | Old | Old
Assay | Old | Old